## Slide 1: Hardware Triggered Scanning: Hardware
Philip Taylor, Emma Arandjelovic
Observatory Sciences Limited

## Slide 2: Overview
T
Motion trajectory
control
Flexible triggering, and fast position capture
Data capture
PandABox
Transformed trigger pulses
Averaged position data
Detector
images
Motor drive
Encoder signal
Raw trigger signals

## Slide 3: Typical Triggering Setup
1
Detector exposing
Detector readout
Trigger at start of each frame
2
3
4

## Slide 4: Triggering Setup: Warning
1
Detector exposing
Detector readout
2
Trigger received when detector still reading out…

## Slide 5: Triggering Setup: Warning
1
Detector exposing
Detector readout
2
Trigger received when detector still reading out…
… leads to missing frames!!

## Slide 6: Trigger Signals
The PMAC has 3 GPIO output signals, generated during a scan
They are fed into the PandA box, which in turn triggers detector acquisition
There are a wide range of detectors in use at Diamond which may require different trigger signals – the PandA can be configured to output an appropriate signal
The GPIO signals are:
  Live: The start of a frame
  Dead: The start of a period when no frames are acquired
  Centre: The middle of the current frame (not normally used)

## Slide 7: Snake Scan

## Slide 8: Motion Control Hardware
Motion control performed by Delta Tau PMAC
Note: Must be CLIPPER or Geobrick LV IMS II
  (old style Geobricks don’t have enough memory)
Co-ordinate system needed to run motion program

## Slide 9: Trajectory Control Overview
Precise control of the trajectory is needed
For each step, the user specifies the end position, the velocity at the end point and the time period for that step
This is called PVT mode:
  Position
  Velocity
  Time

## Slide 10: Trajectory Control Sequence
Malcolm generates trajectory positions and delta times using the Python module scanpointgenerator
EPICS writes them to the PMAC via the pmac driver
The PMAC calculates velocities and performs the scan using PVT mode
PMAC User Subroutines, which output GPIO signals, are called as each scan point is reached. 8 routines are available, corresponding to all combinations of the 3 signals: Live + Dead + Centre

## Slide 11: PandABox Overview
Position and Acquisition Box
Flexible, modular platform
Supports a variety of scanning applications
Based on a PicoZed module which integrates:
  ARM processor running embedded Linux
  FPGA firmware providing configurable function blocks

## Slide 12: PandA Hardware

## Slide 13: Functional Blocks
Each one performs a specific function, e.g. position capture
Configured via a number of registers
Wired up and configured at runtime via a TCP server
Configurations saved to disk

## Slide 14: Available Function Blocks
Function Block | Description
INENC, OUTENC | Input and output encoder signals
TTLIN, TTLOUT | TTL Input and output signals
COUNTER | Up/down pulse counter
CALC | Sums up to 4 position inputs
LUT | Performs logical function on 5 inputs
PCAP | Position capture – options for min, max, sum
PULSE | One-shot pulse delay and stretch
SRGATE | Set Reset Gate – set configurable high or low output
…. | ….
For the full list and more information, see:
https://pandablocks-fpga.readthedocs.io/en/latest/blocks.html

## Slide 15: Configuring PandA
PandA runs an embedded web server
Point a browser directly to this for low level configuration such as updating packages
For normal use, use Malcolm web gui:
  Test rig: http://localhost:8008/gui
  Beamlines: http://ixx-control:8008/gui

## Slide 16: PandA Configuration GUI